## Main
Item: | GJU General Specifications | GJU Technical Specifications
Equipment1: | Industrial Tool Hard Case (Professional Case for Energy Auditing)
Specifications | Heavy-duty protective case for energy auditing and field service tools. Designed for secure transport and organization of disassembly/assembly tool kits and sensitive measurement devices. High-capacity interior with customizable foam compartments and repositionable dividers. Weather-resistant and dust-tight construction (IP67 certified) for rugged field use. Optional: Integrated four-point locking system for security. Features a retractable telescoping handle and in-line wheels for easy mobility. Approx. External Dimensions: 20" × 15" × 8". | . Lightweight design under 6.5 lbs for easy transport. . Waterproof hard-bottom base for tool protection and stability. . Multi main compartments with multi pockets for organized storage.
QTY [Units] | 3
Equipment 2: | Tool Bucket Organizer Bag (Pro-Grade Bag)
Specifications | Pro-grade organizer bag designed to fit standard 5 or 6-gallon buckets. Suitable for energy auditing and general tool organization. High-capacity design with approximately 30 dedicated pockets, sleeves, and loops for tools and accessories. (Multiple Storage compartments) Constructed from durable, abrasion-resistant polyester with a rigid base for protection. Features a high-visibility interior and removable dividers for easy tool identification and customization. Offers weather resistance (IP64 rating) against dust and water splashes. Includes padded carry handles and external webbing for attaching additional gear. | . Durable, waterproof hard-bottom base. . Rugged exterior shell for tool protection. . multi internal pockets for organization.
QTY [Units] | 3
Equipment 3: | Electricity Smart Meter (Single & Three Phase)
Specifications | Advanced smart meter for single and three-phase electrical systems. Provides high-precision measurement of key electrical parameters including voltage, current, power, energy, and power factor. Features multiple communication options including LPWAN (NB-IoT, LTE-M), RF (LoRaWAN), Wi-Fi, and Bluetooth for flexible data transfer. Robust, weatherproof housing (IP65 rated) suitable for harsh environments. | Supports configurable multi-tariff billing and stores over 365 days of load profile data. Includes advanced tamper detection and a high-visibility, backlit display. Designed for easy integration into energy management and billing systems using standard protocols. Certified to international metering standards (MID, ANSI, IEC).
QTY [Units] | 3
Equipment 4: | Electrical Line Splitter (Professional Single-Phase Circuit Splitter & Measurement Interface)
Specifications | Professional tool for safely accessing single-phase electrical circuits for measurement. Non-invasively separates live and neutral conductors to create accessible test points without cutting wires or disrupting power. Enables easy and safe connection of clamp meters, power analyzers, and energy loggers. Rated for continuous use up to 50A AC and 600V AC CAT III. | Constructed with a high-impact, flame-retardant housing and high-conductivity internal components. Features insulated, color-coded 4mm safety jacks and integrated overload protection. Safety certified to international standards (IEC/EN 61010-031, UL 61010-031). Compact, handheld design with a non-slip base for stability.
QTY [Units] | 3
Equipment 5: | Power supply 1 phase /3 phase Ac& Dc
Specifications | 1. Input Power & Compatibility: Input: o 380–400 V AC, 50Hz o 3Ph + N + G (three-phase with neutral and ground) • Variable Outputs: o Three-phase AC output:  Adjustable voltage: 0 to 400 V  Maximum current: 10 A o DC output:  Adjustable voltage: 0 to 230 V  Maximum current: 16 A • Independent voltmeter and ammeter for the output. With high accuracy (resolution) • Analog or digital type meters to allow clear visualization of variations. • Main automatic circuit breaker for general protection. • High-sensitivity residual current device for earth-leakage protection. • Individual magneto-thermal circuit breakers for each output. • Conforms to CE safety and performance standards. • Manufactured under a certified Quality Management System (ISO 9001 or equivalent). Operating Environment:(Operating temperature: 0 to +45C) and (Relative Humidity: <=90%)
QTY [Units] | 1
Equipment 6: | Three-Phase Motor Test Bench with Dynamometer and Torque-Speed Measurement System
Specifications | General Description • The unit shall consist of a complete test bench designed for studying and testing three-phase induction motors under different load conditions. It should include a three-phase motor, an adjustable loading unit, a torque and speed measurement system, and a mechanical mounting/alignment structure that ensures proper coupling and smooth operation of all components. All components must be mechanically and electrically compatible and function together as one integrated test bench. The system should include appropriate electrical and mechanical protections, such as overload and emergency stop functions. | 1. Three-Phase Squirrel Cage Induction Motor • Type: Three-phase squirrel cage induction motor. • Rated power: Within the range of 0.8 to 1.5 kW. • Number of poles: 4 poles (nominal speed between 1350 and 1500 rpm at 50 Hz). • Rated voltage: 380 – 400 V AC, three-phase, 50 Hz. • Rated current: Consistent with the selected power rating. • The motor shall be suitable for connection with a variable load for laboratory experiments. • Mounted on an adjustable base allowing accurate alignment with the load and measuring devices. | 2. Adjustable Loading Unit (Eddy Current or Equivalent Type Brake) • Type: Electromagnetic Eddy Current Brake or an equivalent adjustable electrical loading system. • Absorbed mechanical power: Minimum 0.8 – 1.5 kW. • Load control: Electrically adjustable, manual or via an electronic controller. • The unit shall allow smooth variation of load torque for different test conditions. • Equipped with a suitable cooling system (air or fin-cooled) to prevent overheating during operation. • Capable of continuous operation within the rated power range. | 3. Torque and Speed Measurement System • Type: Integrated torque and speed measurement unit (or two separate sensors). • Torque measurement range: 0 – 10 Nm minimum. • Speed measurement range: 0 – 3000 rpm minimum. • Measurement accuracy: ±1% or better. • Display: Digital display or PC interface (USB or RS232). • Data acquisition capability for performance monitoring and analysis. • Must provide real-time readings of torque and rotational speed during testing. | 4. Mounting and Alignment Structure • The system shall include a mechanical structure that allows secure mounting and precise alignment of the motor, the loading unit, and the measuring device. • The structure should permit linear adjustment of the motor and load to ensure smooth coupling and flexibility in configuration. • It must provide stable and vibration-free operation during tests. • The mounting frame shall be robust, easy to assemble/disassemble, and allow visual access for instructional use. • Equivalent alignment mechanisms (e.g., adjustable rails or sliding platforms) are acceptable, provided they ensure accurate mechanical coupling between the motor and the load.
QTY [Units] | 1
Equipment 7: | Three-Phase Induction Motor Educational Model (Cutaway Type)
Specifications | General Description: An educational cutaway model of a three-phase squirrel cage induction motor, designed to demonstrate the internal construction and components of the motor. The unit is intended for teaching and demonstration purposes only — not for electrical operation. | Technical Specifications: Type: Three-phase squirrel cage induction motor (cutaway / sectioned model). Rated power: Approximately 0.3 – 1.0 kW (similar in size to standard laboratory motors). Voltage: Not intended for electrical connection or operation. Design: Partially cut or transparent sections to clearly display internal components such as: Stator with windings Rotor (squirrel cage type) Bearings Fan Front and rear covers All internal parts must be securely fixed and safely enclosed for close visual inspection and safe handling. The model must be mounted on a stable metal base suitable for classroom or laboratory display. | All internal parts must be securely fixed and safely enclosed for close visual inspection and safe handling. The model must be mounted on a stable metal base suitable for classroom or laboratory display.
QTY [Units] | 1
على ان لا تتجاوز مدة التوريد 3 أسابيع من تاريخ التبلغ