## Main
Item: | GJU General Specifications | GJU Technical Specifications
Equipment 1: | Industrial Power Clamp Meter (True RMS - Advanced Series)
Specifications | 1- Industrial-grade multimeter with clamp functionality. 2- True RMS measurements for accurate readings on non-linear loads. 3- Bluetooth-enabled for wireless data logging via Link app. 4- Designed for energy auditing, electrical maintenance, and power quality analysis. 5- CAT 1000V safety rated for high-energy environments. 6- Complete kit includes carry case, holster, and technical manual. | 1. Measurement Capabilities: a. Voltage AC/DC Voltage Range: 0 to 1200 V Basic Accuracy: ±0.7% + 3 digits Input Impedance: ≥10 MΩ b. Current AC/DC Current (Clamp): 0 to 1500 A True RMS for accurate measurement of non-linear, distorted waveforms Inrush current capture and recording function | c. Power Parameters Active Power (P): Measurement up to 1.5 MW Power Factor (PF) Range: 0.001 to 1.000 Harmonic Analysis: Up to the 63rd order d. Additional Measurements Resistance: 0 to 200 kΩ Capacitance: 0 to 2000 µF Frequency Response: 10 Hz to 2 kHz Temperature: -40°C to 1200°C (with optional probe) | 2. Special Features: a. Phase Sequence / Rotation Indicator b. Non-Contact Voltage Detection (NCV) c. Audible and Visual Continuity Alert (<25 Ω) d. Diode and Transistor Test e. Data Functions Data Hold, Min/Max/Avg recording, Peak capture (1 ms) Relative differential measurement mode 3. Display & Interface: a. High-Resolution LCD with adaptive auto-backlight b. Ultra-Bright white LED area light for low-light conditions c. Wireless Connectivity | 4. Safety & Compliance: a. Safety Certification: CAT IV 1000V / CAT III 1200V b. Insulation: Double-layer reinforced insulation
QTY [Units] | 2
Equipment 2: | Industrial Tool Hard Case (Professional Series for Energy Auditing)
Specifications | 1- Heavy-duty protective case for energy auditing equipment. 2- Specifically designed for disassembly/assembly tool kits. 3- High-capacity storage with customizable compartments. 4- Weather-resistant construction for field use. 5- Secure transport of sensitive measurement devices. | 1. Construction & Materials: a. Exterior Shell * Material: High-strength, co-polymer polypropylene composite * Frame: Heavy-duty, corrosion-resistant aluminum alloy reinforcement * Corners: Triple-thick, over-molded rubber corners for maximum impact absorption b. Interior Protection * Padded foam lining: 25mm thick pick-and-pluck foam with solid foam base layer * Compartment System: Fully customizable, repositionable dividers * Non-slip base: Integrated vibration-dampening rubber feet 2. Approx. Dimensions & Capacity: a. External Dimensions: 20" × 15" × 8" (L×W×H) b. Internal Dimensions: 19" × 14" × 7.5" c. Weight Capacity: 25 kg (55 lbs) | 3. Protective Features: a. Sealing System * Silicone rubber gasket (full perimeter, continuous channel) * Environmental Rating: IP67 certified (dust-tight and protected against temporary immersion in water) b. Security & Locking * Four-point latching system with integrated lock cores (keyed alike) * Automatic pressure equalization valve 4. Organizational Features: Multi-Level Storage: Lower Compartment: Deep, foam-padded cavity for large diagnostic tools Lid Organization: Clear, zippered mesh pockets and elastic retention straps | Removable Multi-Tool Tray: With dedicated molded slots for meters, probes, and accessories 5. Additional Features: a. Ergonomics & Transport * Handle: Contoured, non-slip comfort grip * Retractable telescoping handle and in-line wheels for extended mobility b. Identification & Traceability * Weatherproof label holder on multiple sides * RFID tag pocket for asset tracking (optional)
QTY [Units] | 3
Equipment 3: | Tool Bucket Organizer Bag (Pro-Grade Bag)
Specifications | a- Industrial Equipment Holder - around 6-gallon. b- Suitable for Energy Auditing Purposes. c- High Capacity for Tool Storage. d- Built with High Quality and Durability. e- Specifically Designed for Holding Sensitive Measuring Devices. | 1. Construction & Materials: a. Base & Frame * Rigid, structural polymer base with integrated non-slip pads. * Semi-rigid, impact-resistant shell panels for full perimeter protection. b. Primary Material * Abrasion-resistant 1680D ballistic polyester with PVC backing for maximum durability and weather resistance. 2. Capacity & Organization: a. Pocket System * Around 30 dedicated pockets, sleeves, and loops of various sizes for tools and accessories. * Includes specialized pockets for: screwdrivers, pliers, power drill bits, and smartphones. b. Internal Organization * Removable dividers and adjustable elastic loops to create a custom layout. * A dedicated zippered utility pouch for small, high-value items. | 3. Protective Features: a. Weather Protection * Fully seam-sealed construction and water-repellent coating (IP64 rating against dust and water splashes). b. Security & Containment * Main compartments feature heavy-duty, self-repairing zippers with storm flaps. * Smaller pockets use high-strength hook-and-loop (Velcro-style) closures. 4. Usability & Design: a. Visibility * High-visibility bright orange interior lining for optimal tool identification. b. Ergonomics & Transport * Padded, wide-webbed carrying handle for comfort with heavy loads. * Reinforced grab handles on both sides for two-person lifting. c. Compatibility * Designed to fit all standard 5 or 6 gallons buckets. | 5. Additional Features: a. Hygiene & Maintenance * Anti-microbial coating on interior fabrics to inhibit mold and mildew growth. * All surfaces are wipe-clean and designed for easy emptying of debris. b. Extra Functionality * External MOLLE/PALS webbing for attaching additional gear pouches. * A reinforced D-ring for attaching a carabiner or key leash.
QTY [Units] | 3
Equipment 4: | Infrared & Contact Thermometer (Advanced Dual-Laser Infrared Thermometer with Contact Probe)
Specifications | 1. Measurement Functions: Dual-mode measurement: Infrared (non-contact) and contact probe (K-Type thermocouple). Suitable for HVAC, industrial, and energy auditing applications. Adjustable emissivity for improved accuracy on different surface types. Laser targeting system for quick spot identification. Rugged, portable design for field and industrial use. | 2. Temperature Ranges: Infrared (non-contact): approx. -30°C to 1000°C (or wider, depending on model). Contact probe (K-Type): approx. -30°C to 600°C. Operating ambient temperature: approx. -10°C to 50°C. | 3. Accuracy & Resolution: Typical accuracy: within ±1.5% of reading or ±1.5°C (whichever is greater). Resolution: 0.1°C / 0.1°F. 4. Display & Interface: Clear digital display with backlight for field visibility. Single or dual-laser targeting (depending on model) to indicate measurement area. Easy-to-use interface for switching between IR and contact modes. | 5. Additional Features: Compatible with standard mini-connector K-Type thermocouples. Designed for energy efficiency audits, HVAC troubleshooting, and industrial temperature checks. Durable housing suitable for frequent field use.
QTY [Units] | 2
Equipment 5: | Electricity Smart Meter (Single & Three Phase)
Specifications | 1. Core Measurement & Accuracy: Measurement: High-precision real-time monitoring of voltage (V), current (A), active/reactive/apparent power (kW/kVAR/kVA), energy (kWh/kVARh), frequency (Hz), and power factor. Accuracy Class: Exceeds standard requirements, achieving Class 0.5S for active energy and Class 1.0 for reactive energy per IEC 62053-22. Nominal Inputs: Direct connection up to 125A, or via external current transformers (CTs) with 1A or 5A secondary. 2. Communication & Connectivity: Dual-Network Redundancy: Integrated modular communication bays supporting simultaneous LPWAN (NB-IoT, LTE-M) and RF (LoRaWAN Class A) connectivity. | Advanced Local Interfaces: Integrated high-speed Wi-Fi (802.11 b/g/n/ac) and Bluetooth for on-site commissioning and data interrogation. Optical Port: IrDA-compliant optical port for secure, direct local readout and configuration. External Antenna: Provisions for an external, ruggedized antenna to ensure signal strength in challenging locations. 3. Data Management & Features: Multi-Tariff & Billing: Supports configurable multi-tariff schedules (8+ tariffs) with a integrated real-time clock (RTC) synchronized via NTP (network time protocol) or GPS for absolute time accuracy. Automatic daylight saving time (DST) adjustment. Data Logging: Extensive internal memory for high-resolution load profile logging (e.g., 30- or 60-minute intervals) with storage for 365+ days of data. | Tamper & Security: Advanced tamper detection for magnetic influence, case opening, reverse current, and neutral disconnect. Events are time-stamped and logged. Registers: Comprehensive registers for import/export active and reactive energy in all tariff periods. 4. Environmental & Construction: Robust Enclosure: High-impact-resistant housing with an IP65 rating for complete protection against dust and low-pressure water jets. Extended Operating Range: -40°C to +70°C (-40°F to +158°F) for reliable operation in extreme climatic conditions. Long-Term Reliability: Designed for a minimum service life of 15 years with ultra-low self-consumption. | 5. Compliance & Integration: International Certifications: MID, ANSI C12.1/C12.20, IEC 62052-11/62053-21/22, and compliance with international RF and safety standards. Sealing: Independently, securely sealed meter body and terminal cover to prevent unauthorized access. System Integration: Open API and support for standard protocols (DLMS/COSEM, Modbus) for seamless integration into SCADA, BMS, and any energy management system. 6. Display & Interface: High-Visibility Display: Backlit, 8-digit LCD with multi-screen capability to cycle through key parameters. High-contrast and wide viewing angle.
QTY [Units] | 4
Equipment 6: | Laptop for Presentations & Instructing Purposes (Ruggedized Convertible) | [5 empty] | ….....
Specifications | 1. Design & Build: Convertible 2-in-1 design (laptop + touchscreen tablet mode with pen input). Ruggedized housing with enhanced durability (meets military or high-grade or equivalent standards for shock, dust, and spill resistance). Lightweight yet durable design for field and classroom use. | 2. Performance: Latest-generation multi-core processor (e.g., Intel® Core™ i7 / AMD Ryzen™ 7 or equivalent). Dedicated or integrated graphics processor suitable for engineering software, simulations, and presentations (e.g., NVIDIA® GTX/RTX or AMD Radeon™). Minimum 16 GB RAM, expandable. Fast SSD storage (minimum 512 GB). | 3. Display & Input: 14–15.6 inch touchscreen display, Full HD or higher resolution, with anti-glare coating. Brightness suitable for indoor and semi-outdoor environments (≥ 300 nits). Includes active stylus/pen with palm rejection support. 4. Battery & Power: Long battery life (approx. 8–12 hours depending on workload). Supports fast charging. | 5. Connectivity & Expansion: Multiple ports: USB-C / Thunderbolt, USB-A, HDMI, and audio jack. Latest Wi-Fi 6 / Bluetooth 5 or newer standards. Optional docking support for expanded connectivity. | 6. Security: Biometric login options (fingerprint or facial recognition). Hardware-level security features (TPM or equivalent). 7. Operating System: Professional-grade OS (e.g., Windows Pro or equivalent business-class OS).
QTY [Units] | 1
Equipment 7: | Electrical Insulation Gloves (Professional-Grade)
Specifications | 1. Safety Ratings & Certification: Voltage Protection: Available in multiple classes for maximum application flexibility: Class 00: Maximum Use Voltage 500V AC / Proof Tested to 2,500V AC Class 0: Maximum Use Voltage 1,000V AC / Proof Tested to 5,000V AC Class 2: Maximum Use Voltage 17,000V AC / Proof Tested to 20,000V AC Standards Compliance: Exceeds requirements of IEC 60903, ASTM D120, and NFPA 70E. Certification: CE, UKCA, and OSHA compliant. Independently certified by a Nationally Recognized Testing Laboratory (NRTL). | 2. Material & Construction: Shell Material: Premium, ozone-resistant natural rubber or advanced latex-free dielectric polymer for superior elasticity and dielectric integrity. Interior Lining: Seamless, hypoallergenic nylon or moisture-wicking flock lining for enhanced comfort and reduced fatigue during extended wear. Reinforcement: Integrated, high-abrasion-resistant leather or composite material overgloves (sold separately or as a kit) for mandatory mechanical protection. Cuff Design: Rolled, beaded cuff for secure fit. Extended lengths available (14" to 18") for greater forearm protection. | 3. Durability, Visibility & Ergonomics: Color-Coding: Industry-standard color-coding for instant voltage class identification (e.g., Class 00: Beige, Class 0: Red, Class 2: Yellow). Quality Assurance: 100% electronically tested at high voltage during manufacturing. Each pair is delivered with a individual test certificate. Ergonomic Design: Anatomically pre-curved fingers and textured surface pattern for maximum dexterity, grip, and reduced hand fatigue. | 4. Sizing & Packaging: Comprehensive Sizing: Full range of sizes from 8 to 11 (US sizes), including half-sizes, to ensure a precise and critical fit. Packaging: Sold as a complete pair (left and right hand). Available individually or in multi-pair economy packs. | 5. Storage, Care & Maintenance: Proper Storage: Includes a dedicated storage bag. Must be stored in a cool, dark, dry place away from ozone sources, direct sunlight, and compressing objects. Cleaning Protocol: Clean inside and outside with a mild soap and water solution. A specialized rubber conditioner is recommended for long-term maintenance. Inspection Regimen: Mandatory air inflation test (roll test) before each use. Formal inspection required every six months per OSHA standards.
QTY [Units] | 5
Equipment 8: | U-Value Meter (Advanced Building Envelope Thermal Analyzer)
Specifications | 1. Measurement Capabilities & Performance Measurement of U-Value (thermal transmittance) in building envelopes, typical range: 0.1 to 7.0 W/m²K. Surface and ambient temperature measurement range: approx. -20°C to +80°C. Measurement accuracy: within ±(5% of reading or better). Resolution: at least 0.01 W/m²K. Capable of also calculating related values such as R-Value, surface temperature, and ambient temperature. | 2. Sensor Technology & Probes Heat flux sensor/plate with high sensitivity for stable long-term measurements. At least two surface temperature probes (thin, attachable type) and one ambient temperature sensor. Optional support for relative humidity measurement to assess condensation risk. | 3. Data Management & Connectivity Internal memory for extended data logging of multiple projects. Connectivity options such as USB and/or wireless (Bluetooth or Wi-Fi) for data transfer. Software support for data visualization, trend analysis, and report generation, with export to common formats (CSV, PDF, Excel). | 4. Display & User Interface Integrated display (LCD or equivalent) showing measurement progress and results in real time. User-friendly interface with basic step-by-step measurement guidance. Multi-language support (including at least English, with other languages optional). | 5. Power & Durability Rechargeable battery supporting at least 12 hours of continuous operation. Rugged housing suitable for field use, with minimum IP54 protection. Operating temperature range: approx. -10°C to +50°C. Portable and lightweight design, supplied with a protective carrying case including all probes and accessories.
QTY [Units] | 1
Equipment 9: | Occupancy Data Logger (Advanced Wireless Occupancy & Environmental Data Logger)
Specifications | 1. Sensing & Detection Technology Occupancy detection using PIR (Passive Infrared) sensor for motion/presence monitoring. Optional support for basic people counting (non-identifying, anonymous). Integrated environmental monitoring: capable of logging at least temperature (°C/°F), relative humidity (%RH), ambient light (lux), and optionally CO₂ levels (ppm). | 2. Data Logging, Connectivity & Power Sufficient onboard memory for extended data logging (e.g., ≥100,000 records). Connectivity options may include USB, Bluetooth, Wi-Fi, or Ethernet (depending on model). Support for data export to common formats (CSV, PDF, Excel). Cloud or local software integration for data management and reporting. | 3. Power Management & Durability Battery-powered design with long operational life (months to years depending on interval). Options for USB or external power supply. Enclosure designed for indoor monitoring environments; protection rating typically IP40–IP65 (depending on model). Operating temperature range: approx. -10°C to +50°C. | 4. Data Analysis & Outputs Provides basic occupancy and environmental trend analysis (e.g., time of use, utilization patterns). Compatible with third-party building monitoring/management platforms through data export or available APIs. Graphical and tabular reporting through companion software or cloud dashboards. | 5. Privacy & Security Occupancy detection is anonymous (non-camera based) to ensure user privacy. Standard data protection practices supported by device software/cloud platform.
QTY [Units] | 4
Equipment 10: | Digital Pressure Gauge (High-Precision Digital Barometric Pressure Gauge & Altimeter)
Specifications | 1. Measurement Capabilities & Performance Pressure measurement range: approx. 300 to 1,100 hPa (mbar) or wider. Accuracy: within ±0.5 hPa or ±0.05% of reading (depending on model). Resolution: at least 0.1 hPa. Multi-parameter display: absolute pressure, relative pressure (gauge), and altitude. Selectable units: hPa, mbar, inHg, mmHg, psi, kPa, ft, m. | 2. Sensor & Calibration Equipped with a stable, temperature-compensated pressure sensor suitable for long-term monitoring. Supports user calibration or reference adjustment (e.g., sea-level pressure input for altimeter mode). | 3. Display & User Interface Clear, backlit digital display with simultaneous readouts. Functions: hold, min/max recording, trend or barometric change indication. Optional programmable alarms for high/low limits (pressure or altitude). | 4. Data Logging & Connectivity Internal data storage for logged measurements (typical: several thousand records or more). Connectivity options may include USB and/or Bluetooth for data transfer. Data export in common formats (CSV, Excel, text files), compatible with PC or mobile software. | 5. Power, Protection & Durability: Powered by standard batteries (e.g., AA/AAA) or rechargeable battery; optional external power supply. Typical runtime: 100+ hours depending on settings. Rugged enclosure, at least IP54–IP65 protection against dust and water spray. Operating temperature range: approx. -10°C to +50°C (depending on model). 6. Accessories & Certification: Supplied with protective case and basic accessories (e.g., USB cable if applicable). Delivered with factory calibration certificate; option for ISO/NIST traceable certification.
QTY [Units] | 3
Equipment 11: | Electrical Line Splitter (Professional Single-Phase Circuit Splitter & Measurement Interface)
Specifications | 1. Core Function & Application Purpose: A high-integrity, safety-certified device that non-invasively separates the live (L) and neutral (N) conductors of a single-phase electrical circuit. This creates a safe, accessible air gap for simultaneous measurement of voltage and current without de-energizing the circuit or altering permanent wiring. Primary Use: Enables safe and convenient connection of clamp meters, power analyzers, and energy loggers for load profiling, power quality analysis, energy auditing, and verification of electrical installations. | 2. Design & Construction Enclosure: High-impact, flame-retardant PC-ABS composite housing (UL94 V-0 rated). Internal Bus Bars: Precision-machined, high-conductivity tin-plated copper bars to minimize added circuit resistance and voltage drop. Connector Points: Insulated, color-coded (Brown/Blue), 4mm safety banana jack measurement points for secure probe connection. Integrated Protection: Built-in, replaceable high-energy fuse for overload protection on the voltage measurement ports. | 3. Measurement Compatibility & Specifications Voltage Rating: 600V AC CAT III / 1000V AC CAT II Current Rating: Continuous 50A AC, with peak withstand capacity of 100A for 5 seconds. Frequency Range: 50/60 Hz, suitable for harmonics analysis up to the 3kHz range. Inherent Loss: Adds less than 0.5 mΩ of additional resistance to the circuit. Compatibility: Designed for use with all clamp meters, power quality analyzers, and data loggers equipped with standard 4mm test leads. | 4. Safety Features Safety Certification: Certified to IEC/EN 61010-031, UL 61010-031. Physical Safety: Finger-safe terminals, fully insulated connectors, and a guarded design to prevent accidental contact with live parts during use. Stability: Non-slip, reinforced rubber feet to prevent movement or tipping during measurement. | 5. Operational Specifications Cable Capacity: Accommodates sheathed single-phase cable diameters from 4mm to 20mm. Operating Temperature: -25°C to +60°C Dimensions & Weight: Compact, handheld form factor. Approx. 180mm x 100mm x 45mm, 450g.
QTY [Units] | 3
Equipment 12: | IAQ Data Loggers (Advanced Multi-Parameter Indoor Air Quality (IAQ) Monitoring Station)
Specifications | 1. Monitoring Capabilities & Sensor Performance Multi-parameter monitoring of key IAQ indicators: CO₂: typical range up to ~5,000 ppm with accuracy suitable for indoor air quality applications. Volatile Organic Compounds (TVOCs): ppb–ppm range with appropriate resolution for IAQ analysis. Carbon Monoxide (CO): detection up to ~1,000 ppm for safety monitoring. Particulate Matter (PM1.0, PM2.5, PM10): optical/laser scattering method, up to ~1,000 µg/m³. Environmental parameters: Temperature: typical indoor operating range (approx. -20°C to +60°C). | Relative Humidity: 0–100% RH with standard accuracy for IAQ sensors. Barometric Pressure: approx. 300–1,100 hPa. 2. Data Management, Logging & Connectivity Onboard data logging with large internal memory (tens of thousands of records or more); expandable via external memory options (e.g., SD card). Connectivity options may include Wi-Fi, Bluetooth, Ethernet, or cellular, depending on model. Cloud-enabled dashboard or PC software for real-time monitoring, historical data review, and reporting. Open integration via standard protocols (e.g., MQTT, Modbus, BACnet) for BMS/BAS and IoT platforms. | 3. Design, Durability & Power Enclosure rated at least IP40–IP65 (depending on installation environment). Operating temperature range suitable for indoor and semi-industrial conditions. Power supply options: mains (AC/DC adapter), PoE, or rechargeable battery backup for temporary use during power interruptions. | 4. Alerts, Reporting & Compliance User-configurable visual/audible alerts when thresholds are exceeded. Automated reporting options in common formats (PDF, CSV). Supports reporting aligned with international IAQ guidelines/standards (e.g., WELL, RESET Air, LEED). | 5. Calibration & Maintenance Delivered with factory calibration certificate (ISO or equivalent). Supports periodic re-calibration and/or field calibration as required. Long-life, replaceable sensor modules where applicable.
QTY [Units] | 2
Equipment 13: | Plug Load Analyzer (Advanced Plug-in Power & Energy Analyzer)
Specifications | 1. Core Measurement Capabilities Measures key electrical parameters: Active Power (W), Apparent Power (VA), Reactive Power (VAR), Energy (kWh), Power Factor (PF), Voltage (V), Current (A), and Frequency (Hz). True RMS measurement for accurate readings on linear and non-linear loads. Typical voltage range: 100–250 V AC (single-phase). Current measurement range: up to ~20 A (plug-in loads). Accuracy: within ±1% (typical) for power and energy parameters. Records minimum, maximum, and average values over user-defined intervals. Some models may include harmonic analysis (THD-V, THD-I) and waveform snapshots for power quality insights. | 2. Data Management, Logging & Connectivity Built-in memory for extended logging (from several days to weeks depending on interval). Supports data export via USB, SD card, or wireless options (Bluetooth, Wi-Fi, depending on model). Compatible with PC/mobile software for report generation and data analysis. Export formats include CSV, Excel, and PDF for integration with audit reports and spreadsheets. | 3. User Interface & Alerts Clear digital display (LCD/graphical) showing key parameters in real time. Optional alarms/indicators for over-current, over-voltage, or excessive energy use. May allow cost estimation by inputting local tariff rates. | 4. Safety, Compliance & Design Designed for indoor plug-in monitoring with protection against overload. Safety rating: at least CAT II 300V (some models CAT III depending on design). CE / IEC compliant (other certifications depending on brand). Compact design to fit standard electrical outlets without blocking adjacent sockets. | 5. Power & Portability Powered by mains through plug-in connection, with or without an internal rechargeable battery for short-term portable use. Lightweight, portable housing suitable for field audits and classroom demonstrations. Supplied with a protective case or pouch for transport.
QTY [Units] | 1
Equipment 14: | AC/ DC Power Supply (Programmable Multi-Range AC/DC Power Supply)
Specifications | 1. Input Power & Compatibility: Wide Input Range: 90-264V AC, 47-63 Hz (Auto-ranging), for global operation without manual voltage selection. Power Factor Correction (PFC): Active PFC (>0.99) for efficient operation and reduced harmonic distortion on the input line. 2. Output Performance & Specifications: Output Channels: Single output channel with independent AC and DC generation. DC Output Range: 0 to 360 V DC AC Output Range: 0 to 300 V AC (45 Hz – 1 kHz) Current Range: 0 to 20 A DC / AC RMS Power Rating: 2000 W (3kVA for AC) continuous output power. Load Regulation: <0.01% + 2mV (DC), <0.1% (AC) Ripple & Noise: <2 mVrms (DC Mode) | 3. Control & Operation Modes Operating Modes: Constant Voltage (CV), Constant Current (CC), and Constant Power (CP) modes with automatic crossover. Programmability: Fully programmable voltage, current, frequency (AC), and slew rate via front panel or interface. Memory: Stores up to 100 instrument configurations for rapid recall of frequent test setups. 4. Display & Interface High-Resolution Display: 4.3" full-color TFT LCD displaying simultaneous numerical readouts for V, A, W, and waveform graphics (AC mode). Measurement Precision: 1 mV / 1 mA resolution with 0.05% + 3 mV basic accuracy for DC voltage. Remote Sensing: Dedicated sense terminals to compensate for voltage drop in lead wires, ensuring precise voltage at the load. | 5. Protection & Safety Features Comprehensive Protection: OVP (Over Voltage Protection), OCP (Over Current Protection), OTP (Over Temperature Protection), and SCP (Short Circuit Protection) with latching/non-latching modes. Safety Certification: Certified to IEC 61010-1, UL 61010-1, with CAT II rating. Features an isolated output and a safety-enabled front terminal cover. Cooling: Variable-speed fan with intelligent temperature control for quiet operation and optimal cooling. | 6. Connectivity & Programming (Optional): Remote Interfaces: Standard Ethernet (LXI), USB (USBTMC), and optional GPIB for seamless integration into automated test systems (ATE). Software: Includes PC application for control, data logging, and waveform generation. SCPI command set supported. | 7. Physical Specifications Form Factor: Compact, benchtop design with a reinforced steel chassis. Portability: Integrated carrying handle. Approximate dimensions: 425mm (W) x 177mm (H) x 470mm (D). Weight: ~20 kg.
QTY [Units] | 1
Equipment 15: | Flue Gas Anayzer (Advanced Professional Combustion Analyzer)
Specifications | 1. Measurement Capabilities & Sensor Performance Gas Measurement Parameters: Oxygen (O₂): 0–25% using electrochemical sensor (±0.2% vol typical accuracy). Carbon Monoxide (CO): 0–10,000 ppm using electrochemical sensor (±5% of reading typical). Carbon Dioxide (CO₂): 0–20% using NDIR or calculated from O₂ and fuel type (±0.5% vol typical). Nitrogen Oxides (NO/NOx): Optional sensor module, typically up to 5,000 ppm. Sulfur Dioxide (SO₂): Optional sensor module, typically up to 5,000 ppm. | Pressure & Draft Measurement: Differential pressure / draft measurement: range ±500 Pa or greater, resolution ≤1 Pa. Optional gas pressure sensor for line testing (e.g., natural gas/propane supply). Temperature Measurement: Flue gas temperature: up to 1000–1200 °C (depending on probe). Ambient air temperature: -20 to 100 °C. Accuracy: typically ±1–2 °C. | 2. Fuel & Appliance Compatibility Compatible with multiple fuel types: natural gas, LPG, light and heavy fuel oils, biomass, wood, and coal. Pre-configured testing modes for domestic boilers, industrial burners, heaters, and furnaces. 3. Data Management, Logging & Reporting Internal memory for storing at least hundreds of test results with timestamp and site data. Automated efficiency and combustion calculations: CO/CO₂ ratio, excess air, stack losses, NOx where fitted. Connectivity options may include USB, Bluetooth, and/or Wi-Fi (depending on model). PC/mobile software available for downloading, analyzing, and exporting data (formats: PDF, CSV, Excel). . Optional direct printing to portable printers. | 4. Design & Usability Clear color or high-contrast display, readable in bright light, with intuitive navigation. Rugged housing with dust and splash protection (minimum IP40–IP54). Ergonomic, lightweight design with magnetic back or stand for hands-free operation. Interchangeable probe system for different applications. | 5. Safety, Calibration & Maintenance Automatic protection systems (e.g., pump stop at high CO, condensate/water trap, particle filter). Built-in safety or leak test routines for gas installations. Sensors designed for long life with replaceable modules. Calibration supported using test gases, with traceable certificates available. Meets or exceeds international combustion analyzer standards (e.g., EN 50379, BS 7967). | 6. Power & Accessories Rechargeable battery with typical runtime 8–12+ hours continuous use. Power supply via USB or AC charger. Supplied as a complete kit: analyzer unit, probe, filters, hoses, and carrying case. Optional accessories: particulate filters, printer, calibration adapters.
QTY [Units] | 1
Equipment 16: | Daylight Harvesting Data Logger (Advanced Photometric & Occupancy Data Logger for Daylight Harvesting)
Specifications | 1. Core Measurement Capabilities High-Dynamic-Range Illuminance Sensor: Range: 0.1 to 200,000 lux Accuracy: ±5% of reading or ±2 lux (whichever is greater) Spectral Response: Precisely matched to the CIE standard photopic luminosity function (V(λ)) Cosine Correction: True hemispheric (frog-eye) correction for accurate angular response. Integrated Environmental & Presence Sensing: Occupancy Detection: High-sensitivity, quad-element PIR sensor with configurable timeout and sensitivity settings. Ambient Light Source Detection: Capable of distinguishing between daylight and electric light contributions. Supplementary Sensors: Integrated temperature (-20°C to 70°C) and relative humidity (0-100% RH) monitoring for comprehensive environmental profiling. | 2. Data Management, Logging & Connectivity High-Resolution Logging: Onboard memory for over 2 million timestamped data points, supporting configurable logging intervals from 1 second to 1 hour. Multi-Protocol Wireless Connectivity: Simultaneous support for LoRaWAN (for long-range, low-power campus-wide networks), Wi-Fi 6 (for high-speed local data transfer), and Bluetooth 5.2 (for device configuration and direct mobile app access). Cloud Analytics Platform: Secure web-based dashboard for real-time visualization, automated calculation of Daylight Autonomy (DA), Continuous Daylight Autonomy (cDA), and spatial daylight glare probability analysis. Supports multi-device fleet management. | 3. Smart Building Integration & Control Open Protocol Support: Native output of illuminance and occupancy data via BACnet MS/TP, BACnet/IP, Modbus RTU/TCP, and MQTT for direct integration into any BMS, BAS, or lighting control system. Dimming Curve Calibration: Software tool to assist in calibrating and commissioning the closed-loop response of DALI, 0-10V, and KNX-based dimming systems to the measured daylight levels. Predictive Scheduling: AI-driven analytics to model and recommend optimized lighting schedules based on historical daylight availability and occupancy patterns. | 4. Durability, Power & Design Robust Enclosure: IP66-rated (dust-tight and protected against powerful water jets) housing with a UV-stabilized polycarbonate diffuser. Magnetic base and standard tripod mount for flexible placement. Autonomous Power: Primary power via a high-efficiency, integrated solar panel with a supercapacitor and rechargeable lithium-ion backup system for continuous, maintenance-free operation for over 12 months. | 5. Reporting & Compliance Automated Certification Reporting: Generates pre-formatted reports with all necessary data for LEED v4.1 Daylight Credit EQc (Option 2), BREEAM, WELL Building Standard, and other international green building certifications. Data Export: Exports raw high-resolution data and summary reports to CSV, XLSX, and PDF formats. Calibration: Supplied with an ISO/IEC 17025 accredited, NIST-traceable calibration certificate for the photometric sensor.
QTY [Units] | 2
Equipment 17: | Thermocouple Logger (Multi-Channel Thermocouple Data Logger & Analyzer)
Specifications | 1. Measurement Capabilities & Performance: Wide Thermocouple Compatibility: Supports K, J, T, E, R, S, N, B, and C-type thermocouples with automatic recognition. High-Precision Measurement: Resolution: 0.01°C / 0.01°F Basic Accuracy: ±(0.05% of reading + 0.3°C) for Type K thermocouples Temperature Range: -200°C to +1820°C (-328°F to +3308°F), dependent on probe type. Advanced Display Modes: Real-time Display: Shows all channel readings simultaneously on a graphical display. Differential Analysis: Calculates and displays the difference between any two channels (ΔT). Statistical Tracking: Live tracking and time-stamped recording of MIN, MAX, and AVG for each channel. | 2. Data Logging, Memory & Connectivity High-Capacity Data Logging: Internal memory for over 500,000 timestamped readings per channel. Supports microSD card expansion up to 1TB. Flexible Logging Control: Configurable logging intervals from 100 milliseconds to 24 hours, with options for manual, interval, or alarm-triggered recording. Dual Data Transfer: High-speed USB-C for direct PC connection and Bluetooth 5.2 for wireless data streaming and mobile app control. Real-Time Clock: Integrated battery-backed real-time clock for accurate event time-stamping. | 3. Analysis, Alarms & Functions Programmable Alarm System: Configurable high/low alarms per channel with visual, audible, and relay output options. Data Analysis Suite: On-device graphical trend display and post-download analysis software for generating detailed reports (PDF, CSV, XLSX). Calibration & Offset: Menu-guided calibration mode with user-applicable offset adjustments for probe compensation and system calibration. | 4. Design & Power Ruggedized Enclosure: IP54-rated case for protection against dust and splashes in industrial environments. User-Replaceable Battery: Long-life, rechargeable lithium-ion battery providing 50+ hours of continuous operation on a single charge. Supports pass-through power operation via USB-C. Ergonomic Design: Large, sunlight-readable 4.3" color LCD with an intuitive keypad and protective boot. | 5. Software & Integration Professional PC Software: Includes suite for device configuration, real-time monitoring, advanced data analysis, and automated report generation. Protocol Support: Optional RS-232 or Ethernet output for integration into PLCs and SCADA systems.
QTY [Units] | 2
Equipment 18: | 3-Phase Squirrel Cage Motor Test Bench (Advanced Modular 3-Phase Motor Performance Test System, 1kW)
Specifications | 1. Motor & Load System Test Motor: High-efficiency (IE3/IE4) 3-phase squirrel cage induction motor, 1.1kW nominal, 4-pole. Wye-delta configured for dual voltage operation (220V Δ / 380V Y, 50/60 Hz). Loading Unit: Regenerative, 4-quadrant AC dynamometer with active inverter control. Capable of both motoring and generating modes for full performance mapping. Rated Power: 5 kW continuous. Torque & Speed Measurement: High-precision, rotary torque flange transducer. Rated Torque: 20 Nm. Accuracy: ±0.1% of full scale. Speed Range: 0-6,000 RPM. Modular Base Plate: Precision laser-aligned, rigid steel T-slotted rail system with calibrated laser alignment tools for precise coupling of motor, torque sensor, and dynamometer. | 2. Measurement Capabilities & Accuracy Electrical Analysis: Power Analyzer: 3-phase, 4-channel true power analyzer with 0.1% basic accuracy for active power (W). Parameters: Measures V, A, W, VA, VAR, PF, Hz, and performs harmonic analysis up to the 50th order as standard. Bandwidth: 5 MHz sampling for capturing transient events and inverter-driven waveforms. Mechanical Analysis: Direct Measurement: Torque (Nm), Rotational Speed (RPM), Mechanical Power (kW). Calculated Parameters: Slip, efficiency (real-time, input vs. output), and load percentage. Speed Control: Fully programmable load profiles (constant torque, constant speed, constant power). | 3. Data Acquisition, Control & Software Integrated DAQ System: 24-bit resolution data acquisition card for synchronous sampling of all electrical and mechanical parameters at 100 kS/s. Advanced Control Software: PC-based software for full system control, real-time data visualization, and automated test sequencing. Features include: Performance Mapping: Automated generation of torque-speed curves, efficiency maps (per IEC 60034-2-1), and circle diagrams. Data Logging: Exports comprehensive data sets to CSV, MATLAB, and XML formats. User Interface: Local touchscreen HMI for basic control and monitoring, plus remote desktop operation. | 4. Safety, Protection & Compliance Integrated Safety System: Multi-level protection including overspeed, over-torque, over-current, and over-temperature shutdowns. Hardware-based emergency stop circuit (Category 3, PLd). Electrical Safety: All terminals and connections rated for CAT III 1000V. Fully insulated safety covers and interlock switches. Certifications: CE, UKCA marked. Full ISO 17025 accredited, NIST-traceable calibration certificate for the entire measurement chain (electrical and mechanical) available. | 5. System Configuration & Options Base Configuration: Includes motor, torque sensor, regenerative dynamometer, T-slotted base plate with alignment tools, and full sensor cabling. Essential Options: Variable Frequency Drive (VFD): For testing motors under inverter supply conditions. Thermocouple Input Module: For simultaneous motor winding and bearing temperature monitoring. Vibration Sensor Kit: For correlated vibration analysis during performance testing.
QTY [Units] | 1
Equipment 19: | Induction Motor Cut-Away Model (Squirrel Cage Model): Advanced Functional Cut-Away Model of a 3-Phase Squirrel Cage Induction Motor
Specifications | 1. Design & Construction: Precision Engineering: Professionally machined sectional cuts to expose all critical internal components with maximum clarity, while maintaining structural integrity for operation. Revealed Components: Stator Assembly: Laminated steel core with embedded, color-coded windings (U, V, W phases clearly distinct). Rotor: Full squirrel cage design with clearly visible die-cast aluminum rotor bars and short-circuiting end rings. Full Drive Train: Precision bearings, shaft, and cooling fan assembly. External Features: Terminal box with accessible connection points (star/delta configurable), and external cooling fins. Mounting: Heavy-duty, anti-vibration aluminum base plate with protective rubber feet. Includes mounting points for bench securing. | 2. Functionality & Operation Dual Operational Modes: Manual Operation: Hand-crank for step-by-step visualization of component interaction. Motorized Operation: Low-voltage, variable-speed DC drive system (0-500 RPM) for smooth, controlled demonstration under power. Includes a PWM speed controller. Functional Electricals: All internal circuits are fully functional. The motor can be connected to a 3-phase supply (via a included safety isolation transformer) to demonstrate operation under load with appropriate lab equipment. | 3. Technical Specifications Motor Type: 3-Phase, 4-Pole Squirrel Cage Induction Motor Nominal Rating: 0.5 kW (approx.) Standard Voltages: Model configured for 220/380V Y/Δ operation (50/60 Hz) Construction: Industrial-grade components to ensure durability despite the cut-away sections. | 4. Educational Features Interactive Elements: Key components are numbered and correlate to a comprehensive instructional guide. Integrated Measurement Points: Test points for safely probing voltages and continuity at the terminal box and stator windings. Augmented Reality (AR) Ready: QR codes on the base link to online interactive content, 3D animations, and explanatory videos detailing motor operation and theory. | 5. Safety & Compliance Built-in Safety: All exposed rotating parts are protected by a clear, high-strength polycarbonate safety guard. Educational Compliance: Designed to meet instructional safety standards for vocational and university engineering laboratories. Certification: CE marked; includes a safety instruction manual. | 6. Accessories & Documentation Included Accessories: Safety guard, hand crank, low-voltage DC motor controller, power supply, and a set of connection leads. Comprehensive Documentation: Full set of electrical diagrams, cross-sectional blueprints, a theory of operation manual, and a suggested curriculum for lesson plans.
QTY [Units] | 1
Equipment 20: | 4-channel AC current data logger (Advanced Multi-Channel Wireless AC Current Logger)
Specifications | 1. Sensor Technology & Performance Flexible Rogowski Coils: Four high-accuracy, 24-inch flexible Rogowski current sensors. Max Conductor Size: 8 inches (203 mm). Wide Current Ranges: User-selectable per channel: 0–500 A AC and 0–5000 A AC True RMS. Voltage Input (Optional): Four optional voltage leads allow for true power (kW), energy (kWh), and power factor measurement on each channel. Measurement Performance: Accuracy: ±0.5% of reading ±5 mA (for current) Resolution: 0.001 A (on the 500A range) Bandwidth: 10 Hz – 5 kHz for accurate capture of harmonics and inrush currents. Sample Rate: 100 kS/s per channel, synchronized. | 2. Data Logging & Memory High-Capacity Storage: 16 GB of internal non-volatile memory, capable of storing over 2 years of data at standard aggregation intervals. Advanced Logging Modes: High-Speed Logging: Configurable from 100 ms intervals for transient capture. Aggregated Logging: Min, Max, Average, and Interval recordings from 1 second to 60 minutes. Event-Triggered Logging: Start/stop recording based on user-defined current or power thresholds. | 3. Connectivity, Power & Environmental Tri-Mode Connectivity: Wi-Fi 6: Direct connection to enterprise networks for real-time cloud streaming. Bluetooth 5.2: For local device configuration and data download to a mobile app. Gigabit Ethernet: For high-speed, reliable data transfer in fixed installations. Extended Power Options: Internal Battery: Rechargeable Li-ion battery providing 7+ days of continuous high-speed logging. External Power: USB-PD (Power Delivery) for continuous, unlimited operation. Supports connection to external 12-24V DC sources for permanent installations. | Ruggedized Design: Environmental Rating: IP68 rated (submersible to 1.5m for 30 minutes) and IK08 impact resistance. Operating Temperature: -20°C to +70°C (-4°F to +158°F) 4. Safety, Compliance & Software: Safety Certification: CAT IV 1000V / CAT III 1500V safety rating. Certified to IEC 61010-1, IEC 61010-2-032, and IEC 61326-1 for EMC immunity. Advanced Software Suite: Includes desktop and web-based software for device configuration, real-time monitoring, detailed power quality analysis (including harmonics), and automated report generation in PDF, CSV, and Excel formats. | 5. Accessories & Package: Included Professional Kit: 1-Main Unit: 4-channel data logger. 2-Sensors: (4) 24-inch flexible Rogowski coils. 3- Voltage Leads: (1 set) for power measurement (optional inclusion). 4- Field Transport: Ruggedized, water-resistant hard case with custom foam insert. 5- Cabling: (1) 10 ft USB Type-A to USB-C cable, (1) 6 ft waterproof USB cable. 6- Power: (1) USB-PD wall charger, (1) optional 12V DC cable for permanent install. 7- Connectivity: Ethernet cable. Replacement Parts: High-capacity 5.0 Ah Li-ion replacement battery.
QTY [Units] | 1
على ان لا تتجاوز مدة التوريد 6 أسابيع من تاريخ التبلغ