## Prince Hamzah Hospital Technical Specifications Item Description: Suction Trolley System (OT) Qty. (2)

| #### Bidder #### Offer | #### Min PHH Requirement | Technical Specifications | No. |
| --- | --- | --- | --- |
|  | Please specify manufacturer and country of origin | Manufacturer | 1 |
|  | Please specify model number of the offered equipment | Model Number | 2 |
|  | FDA approval or CE marking or Declaration of conformity | Safety Standard | 3 |
|  | Compact design, heavy duty and easy to clean It should be spacious enough to allow free movement of personnel, trolleys, stretchers, wheel chairs, monitoring devices, portable x-ray machines etc. | Design | 4 |
|  | Suction trolley units allow to perform suction interventions in a rational and effective manner as they include in one suction system all the necessary devices for a broad range of surgical and medical procedures | Application | 5 |
|  | Universal trolley 5 Anti-static casters with brakes | Trolley | 6 |
|  | Two, 2 L with Silicon Ring Lid | Collection Jar | 7 |
|  | Two , 0-1000 mbar | Vacuum regulator | 8 |
|  | Two, included | Safety Jars | 9 |
|  | (600x600x900) ±10% mm | Dimensions (LxWxH) | 10 |
|  | Two outlets | Terminal units for vacuum regulators connection | 11 |
|  | length 3 m, VACUUM ANTI-SQUEEZING HOSE Securely fixed on the trolley | vacuum supply hose | 12 |
|  | (Different standards available), Fits into the vacuum outlet in the operating room | Probe with hose connector | 13 |
|  | Should be according to a seen & tested sample that should comply with above specifications for offered model. | Unit selection criteria | 14 |